PERFORMANCE OF AN AUTOMATED EXTERNAL DEFIBRILLATOR IN A MOVING AMBULANCE VEHICLE

J-G. Yun1, K-W. Jeung2, B-K. Lee2, H-H. Ryu2, H-Y. Lee3, T. Heo1, Y-I. Min1

1Dongkang College, Gwangju, 2Chonnam National University Hospital, Gwangju, 3School of Military Medicine, Daejeon, South Korea

Objectives: The available data suggest that automated external defibrillators (AED) can be safely used in vibration-like moving conditions such as rigid inflatable boats and aircraft environments. However, little literature exists examining their performance in a moving ambulance. The present study was undertaken to determine whether an AED is able to analyse the heart rhythm correctly during ambulance transport.

Methods: An ambulance was driven on paved (20-100 km/h) and unpaved (10 km/h) roads. The performance of two AED devices (CU ER 2, CU Medical Systems Inc., Korea, and Heartstart MRx, Phillips, USA) was determined in a moving ambulance using manikins. Vibration intensity was measured simultaneously with a digital vibrometer. AED performance was then evaluated again on manikins and on a swine model under simulated vibration intensities (0.5-5 m/s2) measured by the vibrometer in the previous phase of the investigation.

Results: The vibration intensity increased with increasing speeds on paved roads (1.98±0.44 m/s2 at 100 km/h). While driving on unpaved roads, it increased to 6.40±1.06 m/s2. Both AED algorithms analysed the heart rhythm correctly under resting state. When tested on pigs, both algorithms showed substantially degraded performances, even at low vibration intensities of 0.5-1 m/s2, which corresponded to vibration intensities while driving on paved roads at 20-60 km/h. This study also showed that electrocardiograms generated on manikins were more resistant to motion artifacts than were the pig electrocardiograms.

Conclusion: Ambulance personnel should consider the possibility of misinterpretation by an AED when this device is used while transporting a patient.