SIMPLIFIED SUBCUTANEOUS MULTI-COIL ARRAY IMPLANTATION FOR HIGH DEFIBRILLATION THRESHOLD

S.A. Kakodkar, M.J. Shih, C. Madias, K. Krishnan, R.G. Trohman

Rush University Medical Center, Chicago, IL, USA

Objective This study was performed to assess the safety and efficacy of defibrillation threshold (DFT) reduction with a simplified method for addition of a three-finger subcutaneous (SQ) array to an implantable cardioverter defibrillator (ICD).

Background: DFT testing at the time of ICD implantation is performed at most medical centers to ensure a 10-J safety margin.  The most effective management of high DFTs involves the addition of a SQ array.

Methods: A total of 1449 consecutive patients received a defibrillator at Rush University Medical Center from July 1, 2002 to September 30, 2010.  Thirty-four patients (2.3%) required SQ array (Guidant Endotak SQ Array XP, Model 0085) implantation for high DFTs.  Instead of making a separate inframammary incision as recommended by the manufacturer, the three-finger SQ array electrodes were inserted vertically and directed posteriorly through the incision created for the device pocket.

Results: A safety margin of greater than or equal to 10-J was achieved in 29 of 34 patients (85%).  Only 2 of the remaining 5 patients required referral for surgical epicardial patch placement.  The only complication was a trochar-related skin perforation. This was repaired without sequelae.  Using this technique, the mean time for implantation of the SQ array was 29.6 minutes.

Conclusion: High DFTs were present in 2.3% of our patient population. This simplified method of three-finger SQ array implantation via a single incision was highly successful in reducing DFTs without complications.  This approach is less cumbersome than making a separate incision for implantation of a three-finger SQ array.