NEW INSIGHT IN TRANSCATHETER VALVE REPLACEMENT : APPROACHING THE ATRIOVENTRICULAR VALVES. AN ANIMAL STUDY

Y Boudjemline1,2 , G Agnoletti1, D Bonnet1,2, D Sidi1,2, P Bonhoeffer3

1Pediatic Cardiology, Necker for Sick Children Hospital, Paris, France, 2EMUI 0016, Necker University, Paris, France, 3Cardiothoracic Unit, Great Ormond Street Hospital, London, UK

Background: Percutaneous valve replacement is limited to the replacement of semilunar valves in selected patients. no data are presently available on percutaneous valve replacement of such valves. We report our experience of percutaneous replacement of the tricuspid valve in animals.  Methods: A newly designed self expandable stent, forming two 40-mm disks separated by a cylinder with a diameter of 18-mm, was designed. A bovine valve was mounted in the central part of the stent. A PTFE membrane was sutured onto the ventricular disk. Eight ewes were, equally, divided into two groups (sacrifice respectively at one hour and one month). Two additional animals were sacrificed after 2.1 month. Results: 9 out of 10 devices were successfully delivered. In one animal, the device was trapped in tricuspid cordae leading to its incomplete opening. A significant paravalvular leak was noticed in one animal from group 2. One insignificant and asymptomatic paravalvular leak was observed in one animal with the longest follow-up. Mean right atrial pressure increased from 5 to 7-mmHg and did not change during the follow-up. At autopsy, examination confirmed the good position of the devices in successfully implanted animals. Conclusions: Implantation of a semi-lunar valve in tricuspid position is possible in ewes through a transcatheter approach. A special device is needed to allow for valve implantation in atrioventricular position. Further improvements in particular in regard of reducing the rate of early stent fractures will, hopefully, make this technique feasible in humans in tricuspid as well as in mitral positions.