IS VDD PACING AN ACCEPTABLE OPTION IN PATIENTS WITH ISOLATED AV BLOCK AND INTACT SINOATRIAL NODE FUNCTION?

P.J. Parwani1, D.R. Frisch2, B.B. Pavri2

1University of Connecticut Health Center, Farmington, CT, USA, 2Thomas Jefferson University Hospital, Philadelphia, PA, USA

Patients with heart block and intact sinoatrial node (SAN) do not need atrial pacing but benefit from AV synchrony. Although routinely treated with DDD pacemakers, an alternate strategy involves a single VDD lead (floating atrial dipole).

Advantages: shorter procedure, lower subclavian vein occlusion/crush, lower cost, but concerns exist about need for atrial pacing or atrial under-sensing. We reviewed all patients with VDD leads (2001-2009).

Selection Criteria: AV block, no SAN dysfunction, sinus rate>60bpm at implantation. Data from implantation and last follow-up analyzed. Development of SAN dysfunction, atrial fibrillation, beta-blocker utilization and addition of atrial lead were recorded. Over 8 years, 37 VDD leads were implanted (Medtronic 5038/St. Jude 1386); 14 male, age=71.3±17(range=22-97)yrs. Mean ejection fraction=60%; 17 patients were on BB. At implantation: mean sinus rate=84±13.7 (range=63-111) bpm; mean P wave via floating dipole=1.8±0.9 (range=0.6-4.2) mV. No complications related to VDD leads. Mean follow-up=568(range=38-2194) days. At last follow-up: mean HR=78.6±14.5 (range=52-108) bpm; mean P wave=1.2±0.9 (range=0.3-3.4) mV. There was no difference in P wave amplitude between patients with>1 vs.<1 yr follow-up (1.2vs.1.2mV,P=NS). 1 patient required addition of atrial lead in<1 month after implant for unrecognized sinus exit-block. 11 patients had paroxysmal AF, 2 patients developed chronic AF (reprogrammed VVI). None required beta-blocker discontinuation, 4 patients had beta-blocker added. Asymptomatic intermittent atrial under-sensing in 4 patients required reprogramming.

Conclusions: VDD PM system is reasonable option for selected patients with isolated AV block and intact SAN function. Addition of atrial lead was required in 1 of 37 pts.

Beta-blocker use was unrestricted. Atrial sensing and AF detection remained adequate.