CLINICAL UTILITY OF CARDIOBIP IN REMOTE DETECTION OF 12-LEAD ECGS IN PATIENTS WITH EXERCISE-INDUCED ISCHEMIA

S. George1, V. Vukcevic2, B. Bojovic3, I. Gussak1, V. Giga2, I. Stankovic2

1NewCardio, Inc., Santa Clara, CA, USA, 2Clinical Center of Serbia, 3Vinca Institute, Belgrade, Serbia

Background: CardioBip records and wirelessly transmits reconstructed 12-lead ECG (12CB). It has 5 contact points: 3 precordial and 2 on the device top that serve as limb leads when touched by index fingers. ECG reconstruction is by patient-specific matrices. To determine whether 12CB could be used to monitor coronary disease (CAD) patients remotely, we compared 12CB to simultaneous 12-lead ECGs (12L) in patients with CAD, before and after exercise treadmill testing (ETT).

Objectives: to determine (1) whether 12CB can accurately reconstruct 12-lead ECGs in patients during recovery phase of ETT; (2) whether 12CB can be used for quantitative evaluation of ST segments in patients with exercise-induced ischemia. Methods: 47 subjects with angiography-documented CAD underwent ETT. 12L and 12CB were recorded prior to and after exercise and read in blinded fashion by 2 experts.

Results: By 12L, 24/47 patients had positive EST (>1 mV horizontal ST depression), and 23/47 were negative. By 12CB, 23/47 were positive (20 true positive [TP], 3 false positive [FP]), and 24/47 were negative (20 TN, 4 FN). Thus, 12CB was noninferior to 12L in detecting ischemia (p=1.0, McNemar’s test). 12CB also demonstrated noninferiority in number of leads detected (p=1.0, McNemar’s test).

Conclusions: 12CB was noninferior to 12L for ischemia detection, and is useful for remote detection and wireless transmission of ischemic events.